## Specifications
Section | Item | Rank | Specifications
Recitals
0.4 | Authors |  | Max Larin, XIMEA, max.larin@ximea.com
| REV 0.5, 28.11.2013 |  | Initail release
Section 1 | Scope
1.1 | What's it all about. |  | 16Mpix direct phosphor imaging camera
1.2 | Code name, Alias |  | MH160MC-KK-FA
1.3 | Customer
1.4 | Customer contacts
Section 2 | Applicable Documents
2.1 | Public
- Information on competitive products
- CE conformity, EN 55022, EN 55024, EN 61326
2.2 | XIMEA Internal |  | - XIMEA Quality System Record
- ISO9001 Quality Handbook
- Conceptual Drawing of hardware platform
- Kodak specification KAI-16000, REVISION 2.0 PS-0013
Section 3 | Requirements
3.1 | General
| What is required? |  | Low noise, low power consumption, compact cooled CCD camera providing direct imaging from P43 scintillator
3.2 | Performance
3.2.1 | Functionality | A
3.2.2 | Capacity | A
3.2.3 | Max. response time after power up | A | <10s
3.2.4 | Full accuracy after power up | A | <300s
3.2.6 | Results
|  | A | - 1394: live image streaming to the PC
|  | A | - 1394: direct control of camera functionality via operating software
|  | A | - 1394: software trigger mode
3.3 | Design
Electronics
| Specific design requirements | A | Long term stability of bonding CCD to FO under regular cooling on/off cycling
|  | A | Camera shall maintain imaging quality within the warranty period, no bubbles and/or exfoliation is allowed
3.4 | Technical Data of Camera
3.4.1 | CCD sensor
3.4.1.1 | Type | B | Kodak KAI016000-AAA-CR-B1, No microlenses, taped glass, Grade 1
| Area |  | 36 x 24mm, 43mm diagonal
| Pixel size |  | 7.4 x 7.4µm
| ADC |  | 14Bit
| Pixel frequency |  | 24.576MHz and 49.152MHz (HS mode)
| Responsivity |  | QE 12% @550nm
3.4.1.2 | Shutter
| internal / external |  | Internal, global, Interline transfer
3.4.1.3 | Resolution
| number of total pixels |  | 16M active pixels
| line number [pixel] |  | 4872 active pixels
| column number [pixel] |  | 3248 active pixels
| Aspect ratio [columns/lines] |  | 4:3
3.4.1.4 | Defects of CCD sensor |  | Inherited from the Kodak specification
| Major dark field defective pixel | A | ≤ 100
| Major bright field defective pixel | A | ≤ 100
| Minor dark field defective pixel | A | ≤ 200
| Cluster defect | A | ≤ 20, N=10, W=3
| Column defects | A | 0
3.4.1.5 | Full resolution refresh rate |  | Measured at 1ms exposure time
| 4904 x 3280 [fps] | A | 1.2
3.4.1.6 | Binning modes refresh rates |  | Measured at 1ms exposure time
| 2452 x 1640 [fps] | A | 2.4
| 1632 x 1092 [fps] | A | 3.4
| 1226 x 820 [fps] | A | 4.4
3.4.1.7 | HS and HS binning modes refresh rates |  | Measured at 1ms exposure time
| 4904 x 3280 [fps] | A | 2.5
| 2452 x 1640 [fps] | A | 4.9
| 1632 x 1092 [fps] | A | 6.2
| 1226 x 820 [fps] | A | 6.2
3.4.1.8 | Partial Readout
| min/max line number [pixel] |  | min height 16 pix - max 4904 pix; increment 4 pix
| min/max column number [pixel] |  | min width 16 pix - max 3280 pix; increment 4 pox
3.4.1.10 | Camera Gain
| at full resolution [e-/count], typical | A | 2.1 ±0.3
3.4.1.11 | Noise
| Readout noise [e-], typical | A | 10
| Dark Current [e-/p/s], @12°C | A | 40
3.4.1.12 | Linearity
| non-linearity over full dynamic range | A | <2% (of full scale) to 95% of full scale
3.4.1.13 | Dynamic Range
| Full well capacity [e-] |  | 30.000
| Digitization [bits] | A | 14
| Pixel Dynamic Range [dB] | A | 65
| SNRmax [dB] | A | 60
3.4.1.14 | AntiBlooming
| AntiBlooming |  | Yes
3.4.1.15 | Conversation factor
| fix / variable |  | variable via hardware gain, controlled via software
3.4.1.16 | Exposure time
| minimum [µs] |  | 100
| maximum [s] |  | 16
| accuracy [µs] |  | 100
3.4.1.17 | Trigger signal
| camera |  | Yes, LVTTL
3.5 | Camera Module Characteristics
3.5.1 | Image Quality
3.5.1.1 | Non-uniformity
gradient
stripe pattern
pattern in dark image
3.5.2 | Optical Components
3.5.2.1 | Optical Interface
| C-Mount |  | n/a
| F-Mount |  | n/a
| other |  | Custom Flange
3.5.2.2 | Filter |  | None
| Type |  | n/a
3.5.2.3 | Optical Cleanliness
CCD Sensor
max dirt particle size in light path [µm²]
max no of particles per image area [no]
within aperture [mm]
Filter
| max dirt particle size in light path [µm²] |  | n/a
| max no of particles per image area [no] |  | n/a
| within aperture [mm] |  | n/a
3.5.3 | Cooling System
| peltier cooler | A | Yes
| peltier power connection | A | internal
| controlling | A | internal
| temperature [°C] | A | +12
| temperature stability [°C] | A | ±0.1
| temperature non-uniformity [°C] | A | ±0.25
| absolute temperature offset [°C] | A | ±0.5
| monitoring | A | Yes
| heatsink / fan / water cooling | A | heatsink
| additional power supply |  | n/a
| Sensor ambient (open, sealed, vacuum, conservation gas) |  | Inert gas
3.5.4 | Electrical Interface
| ESD signal lines protection |  | yes, 2000V HBM
| wrong polarity protection |  | yes, internal protection circuit
| CE compliance |  | EN 55022, EN 55024, EN 61326
| RoHS conformity |  | yes
3.5.4.1 | IEEE1394A Interface
| type |  | IEEE1394A
| speed [Mbit/s] |  | 400Mb/s
| Input voltage |  | 12V DC
| connectors |  | Standard 1394A
| cable |  | Standard
| max cable length [m] |  | 4.5
3.5.4.4 | Temperature/Humidity Sensor interface
| T-sensor on control PCB [°C] |  | yes, TMP106
| CCD T-sensor [°C] |  | yes, PT100
3.6 | Software Interface
| SDK deliverables |  | SDK library, provided as DLL executable, library and C header file
| SDK compatibility |  | Windows XP SP3, Vista SP2 (x32/x64), W7 (x32/x64)
| SDK functionality baseline |  | Standard XIMEA API, documentation http://www.softhard.com/api
SDK functionality extended
3.7 | Physical Characteristic of Camera
3.7.1 | Camera housing
| max length [mm] |  | 63
| max width [mm] |  | 63
| max height [mm] |  | 39.2
3.7.2 | Board(s)
max length, width, height [mm]
3.7.9 | Connectors of Board |  | IEEE1394A
additional power supply for cooling
trigger connector
3.7.10 | CCD compartment |  | Air tight
| Filling |  | Air replaced with Xenon
3.7.11 | Housing
| form |  | Round with square flange
| material |  | Anodized aluminum alloy
| color | A | black/silver
3.7.12 | CCD sensor alignment |  | Measured against housing mounting holes
| tilt [°] |  | < 1
| rotation [mm] |  | < 0.1
| x-/y-plane [mm] |  | < 0.1
| z-axis [mm] |  | < 0.1
3.8 | Physical Interface
3.8.1 | Standard Communication Interface
| type | A | IEEE1394A
| connectors | A | Standard 6pin
| max cable length [m] | A | 4.5
3.8.7 | Mechanical Interface
Camera body
Board outline and mounting holes
3.9 | Operating Characteristics
3.9.1 | Operation
| standard setup [hrs] |  | continuous operation possible within MTBF, no limits
| Power Consumption |  | 6W max
| Power management |  | Active
3.9.2 | Desktop PC configuration
| details |  | CPU > 2GHz; RAM > 1GB recommended; OHCI compliant 1394 adapter
3.10 | Workmanship
| Appearance |  | Professional appearance and functional design
3.11 | Reliability
3.11.1 | MTBF
| time [hrs] | A | 100.000
3.12 | Environment
3.12.1 | Operating
| Temperature |  | +5…+50 °C
| Performance guaranteed temperature |  | +5…+30 °C
| Temperature protection |  | Yes
| Temperature shock rate |  | +5…+50 °C at 1 °C/min
| Altitude |  | Sea level to 1.500 m
| Humidity |  | 0…70% RH, non condensing
| Environment |  | Air
| EMI/EMC |  | see 3.15
3.12.2 | Storage
| Temperature |  | -20…50 °C
| Temperature shock rate |  | -20…50 °C at 1 °C/min
| Altitude |  | to 1.500m
| Humidity |  | 0…70% RH for 2 years, non condensing
Vibration
Shock
Environment
Climate
3.12.3 | Shipment
| Temperature |  | -25…+70 °C
| Temperature shock rate |  | -25…+70 °C at 1 °C/min
| Altitude |  | to 15000m, cosmic ray defects
| Humidity |  | 0…100% RH for 1 month
| Air pressure |  | 100±15kPa
Vibration
Shock
Environment
Climate
3.13 | Transportability
| transport requirements |  | The camera module has to be delivered in a packaging appropriate for high end optical and electronic devices
3.14 | Interchangeability
Maintenance Level 1
Maintenance Level 2
Maintenance Level 3
3.15 | Safety
3.15.1 | Norms, standards |  | EN 55022, EN 55024, EN 61326, EN 60950
3.16 | Identification and product marking
3.16.1 | External Identification plate |  | Serial number of electronic board
3.16.2 | Digital identification |  | Readout of serial number of electronics and CCD sensor
3.17 | Documentation
| Form |  | Final specification of this device shall be documented in the format of this request sheet or additional documents which are linked from this sheet
Section 4 | Verification
4.1 | Classification of inspections |  | All technical parameters will be approved
4.2 | Inspection conditions |  | Standard laboratory test environment
4.3 | Prototype inspection |  | testing of all technical requirements in section 3
4.4 | Incoming inspection for serial production |  | Standard incoming test
Section 5 | Miscellaneous / Additions
Note on B12: Max Larin: The following documents form a part of the specification.
Note on B28: Max Larin: Functionality and its upper and lower performance characteristic. State what the item or system shall do in terms of capacity or function of operation.
Note on B134: Max Larin: There is temperature nonuniformity across the CCD area due to its structure and location of power dissipating nodes. It is a static gradient pattern and guesstimated value is ~0.25°C
Note on B200: Max Larin: This paragraph is intended to indicate as detailed as practicable the standard of workmanship quality that the product must meet to be acceptable. Only proof by visual examination
Note on B203: Max Larin: Reliability requirements shall be stated numerically with confidence levels, if appropriate, in hardware mean time between failures.
Note on B207: Max Larin: Environments that the system or equipment is expected to experience in operation, storage and shipment.
Note on B240: Max Larin: Specify the requirements for the level at which components shall be interchangeable or replaceable.
Note on B245: Max Larin: Specify requirements to preclude or limit hazard to personnel, equipment, or both. "Fail-safe" and emergency operating restrictions shall be included when applicable.
Note on B252: Max Larin: Where applicable, requirements for documenting the design shall be specified in general terms in development specifications.
Note on C1: Max Larin: A = Must, B = important C = nice to have
Note on D65: Max Larin: A group of 2 to “N” contiguous major defective pixels, but no more than “W” adjacent defects horizontally